## MU Calculator
Anritsu MA242x8A/MA243x0A USB Power Sensor Measurement Uncertainty Calculator v1.0
Power Sensor
USB Power Sensor | MA24340A
Sensor VSWR | 1.2101
Device Under Test (DUT)
Measured Power Level (dBm) | 0
Measured Frequency (GHz) | 7
VSWR | 1.2
Measurement Settings
Averaging | 64
Aperture Time (mS) | 20
Signal Type (CW / Digital Mod) | CW
Enhanced Modulation (Enabled / Disabled)
Ambient Temperature (°C) | 30
Sources of Uncertainty
| Uncertainty Type | Value (±%) | Probability Distribution | Divisor | Standard Uncertainty (%)
Linearity and Cal Factor | Type B | 1.5048 | Normal | 2 | 0.7524
Noise | Type B | 1.8413e-05 | Normal | 2 | 9.2065e-06
Zero Set | Type B | 0.000495 | Rectangular | 1.7321 | 0.00028579
Zero Drift | Type B | 0.000414 | Normal | 2 | 0.000207
Digital Modulation Effect | Type B | 0 | Rectangular | 1.7321 | 0
Mismatch between Sensor and DUT | Type B | 1.7358 | Rectangular | 1.7321 | 1.0022
Measurement Uncertainty
Combined Uncertainty (%) | Root_Sum_Squared |  |  |  | 1.2532
Expanded Uncertainty (%) | Coverage Factor (K) | 2 |  | ± | 2.5064
Expanded Uncertainty (dB) |  |  |  | Upper limit | 0.10751
|  |  |  | Lower limit | -0.11024

## BB1_Data (hidden)
Linearity & Cal Factor |  |  |  |  | Spec | Final Spec | Model
Freq1 | Inc? | Freq2 | Inc? | (dB) |  |  | MA24208A | Linear/CF | 0-50C_HP | 0-50C_MP | 0-50C_LP | 20-30C_HP | 20-30C_MP | 20-30C_LP
0.01 | y | 0.05 | y | 0.111 | 0 | 0 | MA24218A | F1 | 0.143 | 0.14 | 0.141 | 0.115 | 0.111 | 0.127
0.05 | n | 2 | y | 0.102 | 0 | 0 | MA24330A | F2 | 0.15 | 0.144 | 0.138 | 0.097 | 0.102 | 0.121
2 | n | 12.4 | y | 0.128 | 0.128 | 0.128 | MA24340A | F3 | 0.152 | 0.134 | 0.132 | 0.1 | 0.128 | 0.135
12.4 | n | 18 | y | 0.114 | 0 | 0 | MA24350A | F4 | 0.137 | 0.127 | 0.168 | 0.115 | 0.114 | 0.139
[8 empty] | Table offset | 5
Lin & CF (dB) |  | 0.128 |  |  |  |  | Signal Type
Lin & CF (%) |  | 2.9912 |  |  |  |  | CW
Noise |  |  |  |  | Spec | Final Spec | Enahnced Modulation
Power1 | Inc? | Power2 | Inc? | (W) |  |  |  | Noise | 0-13C | 13-38C | 38-50C
-70 | y | -19 | n | 1.445e-07 | 0 | 0 |  | R3 | 1.53e-07 | 1.445e-07 | 1.628e-07
-19 | y | 1 | n | 1.195e-05 | 1.195e-05 | 1.195e-05 |  | R2 | 1.425e-05 | 1.195e-05 | 1.423e-05
1 | y | 20 | y | 0.001009 | 0 | 0 | Power Max (dBm) | R1 | 0.00121 | 0.001009 | 0.001277
[7 empty] | 20 | Table offset | 2
2σ Noise Spec (W) without Avg |  | 2.2583e-07 |  |  |  |  | Power Min (dBm)
Power (W) |  | 0.001 |  |  |  |  | -70
Noise Spec (%) |  | 0.022583
Frequency Max (GHz)
40
Zero Set |  |  |  |  | Spec | Final Spec
Power1 | Inc? | Power2 | Inc? | (W) |  |  | Frequency Min (GHz) | Zero Set | 0-13C | 13-38C | 38-50C
-70 | y | -19 | n | 3.323e-10 | 0 | 0 | 0.01 | R3 | 2.7e-10 | 3.323e-10 | 3.128e-10
-19 | y | 1 | n | 3.873e-08 | 3.873e-08 | 3.873e-08 |  | R2 | 4.431e-08 | 3.873e-08 | 3.929e-08
1 | y | 20 | y | 1.07e-06 | 0 | 0 | Aperture Max (mS) | R1 | 1.182e-06 | 1.07e-06 | 1.663e-06
[7 empty] | 1000 | Table offset | 2
Zero Set (W) |  | 3.873e-08 |  |  |  |  | Aperture Min (mS)
Power (W) |  | 0.001 |  |  |  |  | 0.01
Zero Set (%) |  | 0.003873
Temperature Max (°C)
50
Zero Drift |  |  |  |  | Spec | Final Spec
Power1 | Inc? | Power2 | Inc? | (W) |  |  | Temperature Min (°C) | Zero Drift | 0-13C | 13-38C | 38-50C
-70 | y | -19 | n | 3.435e-10 | 0 | 0 | 0 | R3 | 2.587e-10 | 3.435e-10 | 2.842e-10
-19 | y | 1 | n | 4.294e-08 | 4.294e-08 | 4.294e-08 |  | R2 | 3.965e-08 | 4.294e-08 | 3.467e-08
1 | y | 20 | y | 9.955e-07 | 0 | 0 | LP_MP Crossover with Enh Mod Disabled (dBm) | R1 | 1.55e-06 | 9.955e-07 | 1.361e-06
[7 empty] | -16 | Table offset | 2
2σ Zero Drift (W) |  | 4.294e-08 |  |  |  |  | MP_HP Crossover with Enh Mod Disabled (dBm)
Power (W) |  | 0.001 |  |  |  |  | 4
Zero Drift (%) |  | 0.004294
LP_MP Crossover with Enh Mod Enabled (dBm)
-19
Dig Mod Effect |  |  |  |  | Spec | Final Spec
Power1 | Inc? | Power2 | Inc? | (dB) |  |  | MP_HP Crossover with Enh Mod Enabled (dBm)
-70 | y | -19 | n | 0.08 | 0 | 0 | 1
-19 | y | 1 | n | 0.088 | 0.088 | 0.088
1 | y | 20 | y | 0.067 | 0 | 0 | ADC sampling speed (kSps)
140
Dig Mod Eff (dB) |  | 0.088 |  |  |  |  | Averages Max
Dig Mod Eff (%) |  | 2.0469 |  |  |  |  | 65536
VSWR |  |  |  |  | Spec | Final Spec
Freq1 | Inc? | Freq2 | Inc? | (dB)
0.01 | y | 0.15 | y | -22.12 | 0 | 0
0.15 | n | 2 | y | -24.94 | 0 | 0
2 | n | 12.4 | y | -20.08 | -20.08 | -20.08
12.4 | n | 18 | y | -19.08 | 0 | 0
RL Spec (dB) |  | -20.08
VSWR |  | 1.22

## BB2_Data (hidden)
Linearity & Cal Factor |  |  |  |  | Spec | Final Spec
Freq1 | Inc? | Freq2 | Inc? | (dB) |  |  |  | Linear/CF | 0-50C_+15dBm | 0-50C_MP | 0-50C_LP | 30-50C_+15dBm | 30-50C_MP | 30-50C_LP | Freq (GHz) | CF (%)
0.01 | y | 18 | y | 0.12 | 0.12 | 0.12 |  | F1 | 0.18 | 0.15 | 0.15 | 0.14 | 0.12 | 0.12 | 0.01 | 2.3
18 | n | 50 | y | 0.14 | 0 | 0 |  | F2 | 0.21 | 0.17 | 0.17 | 0.17 | 0.14 | 0.14 | 0.05 | 1.5
[15 empty] | 0.1 | 1.5
7 | [5 empty] | 1.5 | [8 empty] | 0.3 | 1.5
[8 empty] | Table offset | 5 | [5 empty] | 0.5 | 1.5
Lin (%) |  | 0.12 | [12 empty] | 1 | 1.5
Lin & CF (%) |  | 1.5048 | [12 empty] | 2 | 1.5
[15 empty] | 3 | 1.5
[15 empty] | 4 | 1.5
Noise |  |  |  |  | Spec | Final Spec | [8 empty] | 5 | 1.5
Power1 | Inc? | Power2 | Inc? | (W) |  |  | Power Max (dBm) | Noise | 0-25C | 25-35C | 35-50C |  |  |  | 6 | 1.5
-70 | y | -20 | y | 4.2266e-08 | 0 | 0 | 20 | R3 | 3.7065e-11 | 3.53e-11 | 3.7065e-11 |  |  |  | 7 | 1.5
-20 | n | 0 | y | 7.7946e-08 | 7.7946e-08 | 7.7946e-08 |  | R2 | 6.8355e-11 | 6.51e-11 | 6.8355e-11 |  |  |  | 8 | 1.5
0 | n | 20 | y | 7.5432e-07 | 7.5432e-07 | 0 | Power Min (dBm) | R1 | 6.615e-10 | 6.3e-10 | 6.615e-10 |  |  |  | 9 | 1.5
[7 empty] | -70 | Table offset | 2 | [5 empty] | 10 | 1.5
[15 empty] | 11 | 1.5
2σ Noise Spec (W) without Avg |  | 1.473e-09 | [12 empty] | 12 | 1.5
Power (W) |  | 0.001 | [12 empty] | 13 | 1.5
Noise Spec (%) |  | 0.0001473 | [12 empty] | 14 | 1.5
[7 empty] | Frequency Min (GHz) | [7 empty] | 15 | 1.5
[7 empty] | 0.01 | [7 empty] | 16 | 1.5
Zero Set |  |  |  |  | Spec | Final Spec | [8 empty] | 17 | 1.5
Power1 | Inc? | Power2 | Inc? | (W) |  |  |  | Zero Set | 0-25C | 25-35C | 35-50C |  |  |  | 18 | 1.5
-70 | y | -20 | y | 9.68e-11 | 0 | 0 |  | R3 | 1.0164e-10 | 9.68e-11 | 1.0164e-10 |  |  |  | 19 | 2.65
-20 | n | 0 | y | 4.95e-09 | 4.95e-09 | 4.95e-09 |  | R2 | 5.1975e-09 | 4.95e-09 | 5.1975e-09 |  |  |  | 20 | 2.86
0 | n | 20 | y | 1.56e-08 | 1.56e-08 | 0 |  | R1 | 1.638e-08 | 1.56e-08 | 1.638e-08 |  |  |  | 21 | 2.71
[8 empty] | Table offset | 2 | [5 empty] | 22 | 2.87
[15 empty] | 23 | 2.91
Zero Set (W) |  | 4.95e-09 | [12 empty] | 24 | 2.71
Power (W) |  | 0.001 | [12 empty] | 25 | 3.09
Zero Set (%) |  | 0.000495 | [12 empty] | 26 | 2.8
[15 empty] | 27 | 3.04
[15 empty] | 28 | 3.19
Zero Drift |  |  |  |  | Spec | Final Spec | [8 empty] | 29 | 2.68
Power1 | Inc? | Power2 | Inc? | (W) |  |  | LP_MP Crossover (dBm) | Zero Drift | 0-25C | 25-35C | 35-50C |  |  |  | 30 | 2.73
-70 | y | -20 | y | 8.9e-11 | 0 | 0 | -20 | R3 | 9.345e-11 | 8.9e-11 | 9.345e-11 |  |  |  | 31 | 2.74
-20 | n | 0 | y | 4.14e-09 | 4.14e-09 | 4.14e-09 |  | R2 | 4.347e-09 | 4.14e-09 | 4.347e-09 |  |  |  | 32 | 2.87
0 | n | 20 | y | 1.72e-08 | 1.72e-08 | 0 | MP_HP Crossover (dBm) | R1 | 1.806e-08 | 1.72e-08 | 1.806e-08 |  |  |  | 33 | 2.84
[7 empty] | 0 | Table offset | 2 | [5 empty] | 34 | 2.96
[15 empty] | 35 | 2.66
2σ Zero Drift (W) |  | 4.14e-09 | [12 empty] | 36 | 3.62
Power (W) |  | 0.001 |  |  |  |  | ADC sampling speed (kSps) | [7 empty] | 37 | 3.18
Zero Drift (%) |  | 0.000414 |  |  |  |  | 140 | [7 empty] | 38 | 5.07
[15 empty] | 39 | 3.52
[15 empty] | 40 | 2.97
VSWR |  |  |  |  | Spec | Final Spec | [8 empty] | 41 | 10.2
Freq1 | Inc? | Freq2 | Inc? | (dB) | [10 empty] | 42 | 10.08
0.01 | y | 0.05 | n | -10.16 | 0 | 0 | [8 empty] | 43 | 9.66
0.05 | y | 0.15 | y | -22.12 | 0 | 0 | [8 empty] | 44 | 9.25
0.15 | n | 2 | y | -28.3 | 0 | 0 | [8 empty] | 45 | 9.05
2 | n | 6 | y | -22.61 | 0 | 0 | [8 empty] | 46 | 8.83
6 | n | 18 | y | -20.44 | -20.44 | -20.44 | [8 empty] | 47 | 8.87
18 | n | 33 | y | -17.95 | 0 | 0 | [8 empty] | 48 | 9.41
33 | n | 40 | y | -14.88 | 0 | 0 | [8 empty] | 49 | 9.65
40 | n | 50 | y | -13.98 | 0 | 0 | [8 empty] | 50 | 10.49
RL Spec (dB) |  | -20.44
VSWR |  | 1.2101